Gauge 2

Calibration Date: 8/13/07

Date Installed: 9/7/2007

Lickstone Bald (~5680 ft)

35.25.505’, 82.58.285’

Way Point 30, Site 6

Visits:

November 13, 2007

Stakes were pulled up but gauge still seemed relatively level – sign was torn off and wood split. Restabilized gauge. Before and after pictures taken. AAs replaced on old logger.

April 4, 2008

1 stake pulled up, but gauge still fairly level. Batteries changed on both

HOBO and datalogger. Left gauge at 5:30 pm.

July 14, 2008

Gauge almost level.

Hobo got data for only 2 days.

Old logger had full record.

Attempted to change battery in Hobo with no effect.  Put in new Hobo (worked) did test tips and relaunched.

Changed AA’s in logger and changed to correct time.

Cleared brush (pictures taken of plants over tops of gauge when we arrived.)

November 26, 2008 ~1130 am

Gauge almost level, we adjusted.

Both HOBO and datalogger had full record.

Cleaned out debris in gauge.

The filter was misshapen (pictures taken) we bent it back to near-original shape and replaced.

Batteries changed in both loggers, test tips done.

March 27, 2009 ~9:30 am

Full data record on both loggers – AAs changed.

Time 15 min. slow on old logger, adjusted.

No releveling necessary. Cleaned out minimal debris in gauge top.

Weather: overcast.

23 June 2009 (~1200)

Full record on both loggers. Batteries changed on both loggers, test tips and relaunch done on HOBO. The computer battery then died and we were unable to complete test tips on the old logger, but the batteries were replaced. Debris was cleaned from gauge top - the gauge was already level. The gauge base was treated with insecticide and some vegetation was cleared.

October 2, 2009

Full record on both loggers, old logger swapped with new. Batter changed on HOBO. No re-leveling necessary. Some debris in funnel.

May 7, 2010

Gauge in good shape – picture by Daniel

Clean gauge, re-level, no battery change. HOBO times okay – EDT (~5 min off)

Directions

At the Waynesville Watershed. Must contact Kyle and Hugh (watertreatment@townofwaynesville.org and parrot_l@bellsouth.net, respectively) to be escorted up to the gauges. Kyle’s number: 456-8497.

To the watershed: 40W to Exit 27 (74W)

74W to Exit 98 (West Waynesville)

Left off the exit (Hyatt Creek Road)

Left at S Main Street/US-23-BR

Right at Allens Creek Road

Follow the road right when you see the quarry in front of you. Make a left just after the quarry, and follow that road until you reach the watershed on your left. Park there and Hugh will take you up to the gauges.

After parking at the end of the trail, walk up the hill following the fence (wire) which marks the town line. At the top of the ridge go to your right. Walk through the field of blueberry bushes and the gauge is in the middle.