Gauge 10

Calibration Date: 8/28, 8/29, 8/31, 9/1/07

Date Installed: 9/22/07

Beaty Spring Knob, ~4849 ft

Way Point 34, Site 7

35.27.381’, 82.56.806’

Visits:

December 29, 2007

Data downloaded

Gauge condition looks good

Good data from both HOBO and datalogger

Swap out AA batteries from old datalogger

Changed time on old logger to local time

April 27, 2008, ~11:30 am EDT

Data downloaded, good data from both loggers

Changed batteries on both loggers

Changed time to EDT (was set to EST)

Cable connection on old logger is working

Gauge in good overall condition

July 16, 2008, ~12 noon local

Gauge tipped over.

Data record for both HOBO and old logger goes through May 14.

AAs changed on old logger.

Resecured, releveled, brush cleared

Before and after pictures taken.

November 22, 2008, ~10:30 am local

Gauge upright – leaves and debris cleared from top.

Full data record on old logger and HOBO.

Batteries changed and test tips completed on both loggers.

Gauge releveled, sign replaced.

April 18, 2009, ~2:00 pm local

Gauge upright and nearly level, minor readjustments made.

Debris cleared from gauge top (not much).

Full data record on both loggers.

AAs changed in old logger – also had to change cord from logger to rain gauge as they would not communicate for test tips.

Time on old logger about 15 minutes slow, adjusted.

Weather: sunny, calm.

18 July 2009, ~0930

Gauge tipped over, stake pulled out of ground. Data on HOBO and old logger go until 17 May 2009. Cleared brush from the area, and reset the gauge. Replaced batteries in both loggers. Found some ants beneath the cover where the loggers connect to the rain gauge. Cleared these out, treated gauge base with insecticide, and taped a Bounce sheet inside the gauge. Cleared debris from gauge top.

October 3, 2009

~9:30 am

Gauge upright. Leaves covering top. Placed new upslope post to secur. Too rocky for other two. Replaced logger with new. Releveled and cleaned. Weather sunny with rising fog. New battery in HOBO.

May 12, 2010

~9:15 am (overcast)

Gauge upright, some debris in funnel, replaced HOBO battery. Significant re-leveling needed; changed day from 13 May 2010 to 12 May 2010!! Tree leaning downslope.

Directions:

40W to Exit 27  (74)

74W to Exit 98

Left on Hyatt Bridge

Left on South Main

Right at Allens Creek

Left on Camp Branch Road

Left at Wolf Pen Road

Enter Crestview Point. Drive up the main road, veer to left at the gravel road. At the next Y in the road, follow it to the right (up). Straight ahead you can see a trail. If you have a 4WD vehicle you can follow the road until it ends, and then walk up the hill. Once you are on the ridgeline use the GPS to help find the gauge.