## Calib Template
Field Protocol for Rain Gauge Calibration
Station | Date | Time (CST) | Nozzle | Collected | Remarks | Operator:
|  |  | mm/hr | Water (ml) |  |  | Page | 1 of 2
50
50
50
50
50
100
100
100
100
100
100
100
100
100
100
Field Protocol for Rain Gauge Calibration
Station | Date | Time (CST) | Nozzle | Collected | Remarks | Operator:
|  |  | mm/hr | Water (ml) |  |  | Page | 2 of 2
300
300
300
300
300
300
300
300
300
300

## Record. Vol. Templ.
Test of Volume of Mariott Tube SJB-1 Nominally 653 ml
[8 empty] | Operator: | S Burges
Date | Collector | Nozzle size | Number of | Collector | Collected | Density | Volume
| Weight | mm/hr | Mariott | Plus | Water | Grams/ml | Delivered
| grams |  | Tube | Mariott | Temperature |  | ml
|  |  | Volumes | Tube Volume | Deg C
Grams
2004-07-10 | 62.9 | 100 | 1 | 715.8 | 21.8 | 0.99782 | 654.3264
| 715.8 | 200 | 1 | 1368.7 | 21.8 | 0.99782 | 654.3264
| 62.9 | 200 | 1 | 715.5 | 21.8 | 0.99782 | 654.0258
| 715.5 | 200 | 1 | 1367.3 | 21.8 | 0.99782 | 653.224
2004-07-13 | 62.8 | 100 | 1 | 717.8 | 21.7 | 0.99784 | 656.4179 | Tube Temp
| 63.7 | 100 | 1 | 716.5 | 23.3 | 0.99746 | 654.4623 | Tube Temp
| 63.5 | 50 | 1 | 716.3 | 23.5 | 0.99742 | 654.4886 | Tube Temp - possible slight water loss
[5 empty] | 24 |  |  | Tub Temp
| 63.6 | 50 | 1 | 717.2 | 24.5 | 0.99717 | 655.4549 | Tube Temp
| 63.5 | 50 | 1 | 716.6 | 24.9 | 0.99707 | 655.0192 | Tube Temp
| 63.7 | 100 | 1 | 716.8 | 24.9 | 0.99707 | 655.0192 | Tube Temp
Average for 200 mm/hr = 654 ml
Average for 100 mm/hr = 655 ml
Average for 50 mm/hr = 655 ml
Test of Volume of Mariott Tube SJB-2 Nominally 653 ml
[8 empty] | Operator: | S Burges
Date | Collector | Nozzle size | Number of | Collector | Collected | Density | Volume
| Weight | mm/hr | Mariott | Plus | Water | Grams/ml | Delivered
| grams |  | Tube | Mariott | Temperature |  | ml
|  |  | Volumes | Tube Volume | Deg C
Grams
2004-07-12 | 62.8 | 100 | 1 | 717.9 | 23.3 | 0.99746 | 656.7682
| 717.9 | 200 | 1 | 1372.1 | 22.8 | 0.99759 | 655.7804
| 64.7 | 200 | 1 | 718.1 | 21.8 | 0.99782 | 654.8275
| 718.1 | 200 | 1 | 1371.2 | 21.7 | 0.99784 | 654.5137
| 64.7 | 100 | 1 | 718.7 | 22.2 | 0.99773 | 655.488 | Temp not measured - estimated
| 63.8 | 100 | 1 | 718.4 | 22.2 | 0.99773 | 656.0893
| 65 | 100 | 1 | 718.9 | 21.3 | 0.99793 | 655.2564
| 65 | 100 | 1 | 718.9 | 22.2 | 0.99773 | 655.3877 | Repeat for receiving Tub higher T
Nozzle 2 | 63.6 | 50 | 1 | 719.1 | 22.4 | 0.99768 | 657.0243
| 63.6 | 50 | 1 | 719.1 | 23 | 0.99754 | 657.1165 | (55 minutes at room temp)
Nozzle 2 | 63.6 | 50 | 1 | 718.8 | 21.8 | 0.99782 | 656.6315
| 63.6 | 50 | 1 | 718.8 | 22.3 | 0.9977 | 656.7104 | Repeat for Tub Temp
Nozzle 2 | 63.9 | 50 | 1 | 718.9 | 22.1 | 0.99775 | 656.4771
| 63.9 | 50 | 1 | 718.9 | 22.7 | 0.99761 | 656.5692 | Repeat for Tub Temp
Average for 200 mm/hr = 655 ml
Average for 100 mm/hr = 656.5 ml
Average for 50 mm/hr = 656.7 ml
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST |  | Time CST | Collected | Collected | Remarks
|  | Collector |  | Collector | Water (ml) | Water (in.)
|  | Placed |  | Emptied
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST |  | Time CST | Collected | Collected | Remarks
|  | Collector |  | Collector | Water (ml) | Water (in.)
|  | Placed |  | Emptied

## Mariott Tube
Calibration of TB3 Gauge #
Calibration Tube sjb-
Date | Time | Nozzle | Cumulative | Rain Depth | Remarks | Operator:
|  | mm/hr | Tips | per tip |  |  | Page | 1
|  |  |  | mm | Nominal Volume = 653 ml;
|  | 50 |  |  | Calibrated Volume SJB - 1: 655 ml
50
|  | 50 |  |  | Calibrated Volume SJB - 2 656.7 ml
50
50
Average
100
|  | 100 |  |  | Calibrated Volume SJB - 1: 655 ml
100
|  | 100 |  |  | Calibrated Volume SJB - 2 656.5 ml
100
Average
200
|  | 200 |  |  | Calibrated Volume SJB - 1: 654 ml
200
|  | 200 |  |  | Calibrated Volume SJB - 2 655 ml
200
Average
300
|  | 300 |  |  | 200 mm TB3: Volume(ml)*.0318 = rain depth (mm)
|  | 300 |  |  | Depth/Mariott volume = 654 ml = 20.8 mm
|  | 300 |  |  | # Mariott Tubes = 2 m rain = 96
|  |  | Pre Test | 50 lots x 2 MTs
Volume

## TB3 -Calib Template
Calibration of Hydrologic Services TB3 Rain Gauge Number XX -- Serial Number - YY
Calibration Tube Used sjb- 2 (or Duke 1 or Duke 2)
Date | Start | Nozzle | Cumulative | Cum | Depth | Depth | Tips | Remarks | Operator:
| Time | mm/hr | Tips | Vol ml | per tip | per tip | per |  |  | Page 1
2007-06-14 |  |  |  |  | mm | inches | inch | Nominal Volume = 653 ml;
152
|  | 100 | 229
|  | 100 | 303
|  | 100 | 378 |  |  |  |  | Calibrated Volume SJB - 2 656.5 ml
|  | 100 | 454
|  | 100 | 530
100
100
100
100
100
Tips |  |  | 378 | 3272 | 0.27553 | 0.010848 | 92.1854
531
|  | 50 | 606
|  | 50 | 683
|  | 50 | 757
|  | 50 | 829 |  |  |  |  | Calibrated Volume SJB - 2 656.7 ml
|  | 50 | 903
Tips |  |  | 372 | 3275 | 0.28023 | 0.011033 | 90.639
3275
904
|  | 300 | 975
|  | 300 | 1046
|  | 300 | 1117
|  | 300 | 1187 |  |  |  |  | Calibrated Volume SJB - 2 655 ml
|  | 300 | 1260
300
300
300
300
300
Tips |  |  | 357 | 3270 | 0.29156 | 0.011479 | 87.1172
Notes
200 mm TB3: Collected or delivered rain Volume(ml)*.03183 = rain depth (mm)
Depth = | Mariott Volume * |  | 0.031831
Area = | 31415.9265 | square mm
Depth/Mariott volume (approx 653 ml) = |  |  | 20.7856 | mm
Depth/Mariott volume Duke 1 = 653 ml = |  |  | 20.7856 | mm
Depth/Mariott volume Duke 2 = 653 ml = |  |  | 20.7856 | mm

## TB3 1-Calib
Calibration of Hydrologic Services TB3 Rain Gauge Number 1 -- Serial Number -
Calibration Tube Used sjb- 2
Date | Start | Nozzle | Cumulative | Cum | Depth | Depth | Tips | Remarks | Operator:
| Time | mm/hr | Tips | Vol ml | per tip | per tip | per |  |  | Page 1
2007-06-14 |  |  |  |  | mm | inches | inch | Nominal Volume = 653 ml;
5
100
100
|  | 100 | [5 empty] | Calibrated Volume SJB - 2 | 656.6 | ml
100
|  | 100 | 515
Tips |  |  | 510 | 3283 | 0.2049 | 0.0080671 | 123.9604
515
50
50
50
|  | 50 | [5 empty] | Calibrated Volume SJB - 2 | 656.7 | ml
|  | 50 | 1020
Tips |  |  | 505 | 3283.5 | 0.20696 | 0.0081482 | 122.7264
1020
300
300
300
|  | 300 | [5 empty] | Calibrated Volume SJB - 2 | 655 | ml
300
|  | 300 | 1647
Tips |  |  | 627 | 3930 | 0.19951 | 0.0078549 | 127.3088
Notes
200 mm TB3: Collected or delivered rain Volume(ml)*.03183 = rain depth (mm)
Depth = | Mariott Volume * |  | 0.031831
Area = | 31415.9265 | square mm
Depth/Mariott volume (approx 653 ml) = |  |  | 20.7856 | mm
Depth/Mariott volume Duke 1 = 653 ml = |  |  | 20.7856 | mm
Depth/Mariott volume Duke 2 = 653 ml = |  |  | 20.7856 | mm

## TB3 2-Calib
Calibration of Hydrologic Services TB3 Rain Gauge Number 2 -- Serial Number -
Calibration Tube Used sjb-1 (or 2)
Date | Start | Nozzle | Cumulative | Cum | Depth | Depth | Tips | Remarks | Operator:
| Time | mm/hr | Tips | Vol ml | per tip | per tip | per |  |  | Page 1
2004-07-18 |  |  |  |  | mm | inches | inch | Nominal Volume = 653 ml;
152
| 14:17 | 50 | 229 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 14:50 | 50 | 303
| 15:18 | 50 | 378 |  |  |  |  | Calibrated Volume SJB - 2 656.7 ml
| 15:44 | 50 | 454
| 16:14 | 50 | 530
Tips |  |  | 378 | 3272 | 0.27553 | 0.010848 | 92.1854
1 tip to empty 3 ml collected - cum = 5*655 - 3 ml |  |  |  | 3272
531
| 16:51 | 100 | 606
| 17:04 | 100 | 683 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 17:18 | 100 | 757
| 17:32 | 100 | 829 |  |  |  |  | Calibrated Volume SJB - 2 656.5 ml
| 17:46 | 100 | 903
Tips |  |  | 372 | 3275 | 0.28023 | 0.011033 | 90.639
1 tip while rim removed. Cum = 5*655 ml |  |  |  | 3275
904
| 18:02 | 200 | 975
| 18:10 | 200 | 1046 |  |  |  |  | Calibrated Volume SJB - 1: 654 ml
| 18:19 | 200 | 1117
| 18:28 | 200 | 1187 |  |  |  |  | Calibrated Volume SJB - 2 655 ml
| 18:37 | 200 | 1260
Tips |  |  | 357 | 3270 | 0.29156 | 0.011479 | 87.1172
1 tip -19:06- add 1 tip to cum tips cum = 5*654 ml |  |  |  | 3270
Notes
200 mm TB3: Collected or delivered rain Volume(ml)*.03183 = rain depth (mm)
Depth = | Mariott Volume * |  | 0.031831
Area = | 31415.9265 | square mm
Depth/Mariott volume (approx 654 ml) = |  |  | 20.8175 | mm

## TB3 3-Calib
Calibration of Hydrologic Services TB3 Rain Gauge Number 3 -- Serial Number -
Calibration Tube Used sjb-1 (or 2)
Date | Start | Nozzle | Cumulative | Cum | Depth | Depth | Tips | Remarks | Operator:
| Time | mm/hr | Tips | Vol ml | per tip | per tip | per |  |  | Page 1
2004-07-18 |  |  |  |  | mm | inches | inch | Nominal Volume = 653 ml;
152
| 14:17 | 50 | 229 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 14:50 | 50 | 303
| 15:18 | 50 | 378 |  |  |  |  | Calibrated Volume SJB - 2 656.7 ml
| 15:44 | 50 | 454
| 16:14 | 50 | 530
Tips |  |  | 378 | 3272 | 0.27553 | 0.010848 | 92.1854
1 tip to empty 3 ml collected - cum = 5*655 - 3 ml |  |  |  | 3272
531
| 16:51 | 100 | 606
| 17:04 | 100 | 683 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 17:18 | 100 | 757
| 17:32 | 100 | 829 |  |  |  |  | Calibrated Volume SJB - 2 656.5 ml
| 17:46 | 100 | 903
Tips |  |  | 372 | 3275 | 0.28023 | 0.011033 | 90.639
1 tip while rim removed. Cum = 5*655 ml |  |  |  | 3275
904
| 18:02 | 200 | 975
| 18:10 | 200 | 1046 |  |  |  |  | Calibrated Volume SJB - 1: 654 ml
| 18:19 | 200 | 1117
| 18:28 | 200 | 1187 |  |  |  |  | Calibrated Volume SJB - 2 655 ml
| 18:37 | 200 | 1260
Tips |  |  | 357 | 3270 | 0.29156 | 0.011479 | 87.1172
1 tip -19:06- add 1 tip to cum tips cum = 5*654 ml |  |  |  | 3270
Notes
200 mm TB3: Collected or delivered rain Volume(ml)*.03183 = rain depth (mm)
Depth = | Mariott Volume * |  | 0.031831
Area = | 31415.9265 | square mm
Depth/Mariott volume (approx 654 ml) = |  |  | 20.8175 | mm

## TB3 4-Calib
Calibration of Hydrologic Services TB3 Rain Gauge Number 4 -- Serial Number -
Calibration Tube Used sjb-1 (or 2)
Date | Start | Nozzle | Cumulative | Cum | Depth | Depth | Tips | Remarks | Operator:
| Time | mm/hr | Tips | Vol ml | per tip | per tip | per |  |  | Page 1
2004-07-18 |  |  |  |  | mm | inches | inch | Nominal Volume = 653 ml;
152
| 14:17 | 50 | 229 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 14:50 | 50 | 303
| 15:18 | 50 | 378 |  |  |  |  | Calibrated Volume SJB - 2 656.7 ml
| 15:44 | 50 | 454
| 16:14 | 50 | 530
Tips |  |  | 378 | 3272 | 0.27553 | 0.010848 | 92.1854
1 tip to empty 3 ml collected - cum = 5*655 - 3 ml |  |  |  | 3272
531
| 16:51 | 100 | 606
| 17:04 | 100 | 683 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 17:18 | 100 | 757
| 17:32 | 100 | 829 |  |  |  |  | Calibrated Volume SJB - 2 656.5 ml
| 17:46 | 100 | 903
Tips |  |  | 372 | 3275 | 0.28023 | 0.011033 | 90.639
1 tip while rim removed. Cum = 5*655 ml |  |  |  | 3275
904
| 18:02 | 200 | 975
| 18:10 | 200 | 1046 |  |  |  |  | Calibrated Volume SJB - 1: 654 ml
| 18:19 | 200 | 1117
| 18:28 | 200 | 1187 |  |  |  |  | Calibrated Volume SJB - 2 655 ml
| 18:37 | 200 | 1260
Tips |  |  | 357 | 3270 | 0.29156 | 0.011479 | 87.1172
1 tip -19:06- add 1 tip to cum tips cum = 5*654 ml |  |  |  | 3270
Notes
200 mm TB3: Collected or delivered rain Volume(ml)*.03183 = rain depth (mm)
Depth = | Mariott Volume * |  | 0.031831
Area = | 31415.9265 | square mm
Depth/Mariott volume (approx 654 ml) = |  |  | 20.8175 | mm

## TB3 5-Calib
Calibration of Hydrologic Services TB3 Rain Gauge Number 5 -- Serial Number -
Calibration Tube Used sjb-1 (or 2)
Date | Start | Nozzle | Cumulative | Cum | Depth | Depth | Tips | Remarks | Operator:
| Time | mm/hr | Tips | Vol ml | per tip | per tip | per |  |  | Page 1
2004-07-18 |  |  |  |  | mm | inches | inch | Nominal Volume = 653 ml;
152
| 14:17 | 50 | 229 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 14:50 | 50 | 303
| 15:18 | 50 | 378 |  |  |  |  | Calibrated Volume SJB - 2 656.7 ml
| 15:44 | 50 | 454
| 16:14 | 50 | 530
Tips |  |  | 378 | 3272 | 0.27553 | 0.010848 | 92.1854
1 tip to empty 3 ml collected - cum = 5*655 - 3 ml |  |  |  | 3272
531
| 16:51 | 100 | 606
| 17:04 | 100 | 683 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 17:18 | 100 | 757
| 17:32 | 100 | 829 |  |  |  |  | Calibrated Volume SJB - 2 656.5 ml
| 17:46 | 100 | 903
Tips |  |  | 372 | 3275 | 0.28023 | 0.011033 | 90.639
1 tip while rim removed. Cum = 5*655 ml |  |  |  | 3275
904
| 18:02 | 200 | 975
| 18:10 | 200 | 1046 |  |  |  |  | Calibrated Volume SJB - 1: 654 ml
| 18:19 | 200 | 1117
| 18:28 | 200 | 1187 |  |  |  |  | Calibrated Volume SJB - 2 655 ml
| 18:37 | 200 | 1260
Tips |  |  | 357 | 3270 | 0.29156 | 0.011479 | 87.1172
1 tip -19:06- add 1 tip to cum tips cum = 5*654 ml |  |  |  | 3270
Notes
200 mm TB3: Collected or delivered rain Volume(ml)*.03183 = rain depth (mm)
Depth = | Mariott Volume * |  | 0.031831
Area = | 31415.9265 | square mm
Depth/Mariott volume (approx 654 ml) = |  |  | 20.8175 | mm

## TB3 6-Calib
Calibration of Hydrologic Services TB3 Rain Gauge Number 6 -- Serial Number -
Calibration Tube Used sjb-1 (or 2)
Date | Start | Nozzle | Cumulative | Cum | Depth | Depth | Tips | Remarks | Operator:
| Time | mm/hr | Tips | Vol ml | per tip | per tip | per |  |  | Page 1
2004-07-18 |  |  |  |  | mm | inches | inch | Nominal Volume = 653 ml;
152
| 14:17 | 50 | 229 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 14:50 | 50 | 303
| 15:18 | 50 | 378 |  |  |  |  | Calibrated Volume SJB - 2 656.7 ml
| 15:44 | 50 | 454
| 16:14 | 50 | 530
Tips |  |  | 378 | 3272 | 0.27553 | 0.010848 | 92.1854
1 tip to empty 3 ml collected - cum = 5*655 - 3 ml |  |  |  | 3272
531
| 16:51 | 100 | 606
| 17:04 | 100 | 683 |  |  |  |  | Calibrated Volume SJB - 1: 655 ml
| 17:18 | 100 | 757
| 17:32 | 100 | 829 |  |  |  |  | Calibrated Volume SJB - 2 656.5 ml
| 17:46 | 100 | 903
Tips |  |  | 372 | 3275 | 0.28023 | 0.011033 | 90.639
1 tip while rim removed. Cum = 5*655 ml |  |  |  | 3275
904
| 18:02 | 200 | 975
| 18:10 | 200 | 1046 |  |  |  |  | Calibrated Volume SJB - 1: 654 ml
| 18:19 | 200 | 1117
| 18:28 | 200 | 1187 |  |  |  |  | Calibrated Volume SJB - 2 655 ml
| 18:37 | 200 | 1260
Tips |  |  | 357 | 3270 | 0.29156 | 0.011479 | 87.1172
1 tip -19:06- add 1 tip to cum tips cum = 5*654 ml |  |  |  | 3270
Notes
200 mm TB3: Collected or delivered rain Volume(ml)*.03183 = rain depth (mm)
Depth = | Mariott Volume * |  | 0.031831
Area = | 31415.9265 | square mm
Depth/Mariott volume (approx 654 ml) = |  |  | 20.8175 | mm

## Sheet7
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied

## Sheet6
Summary of Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Field Protocol for Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Field Protocol for Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Field Protocol for Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Field Protocol for Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied
Field Protocol for Rain Gauge Volume Collection
Operator:
Station | Date | Time CST | Time CST | Collected | Collected | Remarks
|  | Collector | Collector | Water (ml) | Water (in.) |  |  | Page #
|  | Placed | Emptied