Gauge 1

Calibration Date: 6/24,6/27/07

Date Installed: June 30, 2007

Camp Daniel Boone (~3794 feet)

Near Deep Gap (35.23’53.9, 82.54’46.9)

Visits:

October 16, 2007 12:00 pm

Replaced all 6 batteries of data logger. Changed data logger from 00:00:24 to 12:06:00.

March 5, 2008

Gauge looks to be in good condition – full data record for HOBO – old logger end 14 Feb 08 (batteries dead?) – time on old logger 13:59 EDT –

AAs changed on old logger – battery changed on HOBO

June 4, 2008

Gauge has some dents but is upright and needs only slight leveling.

Full record for HOBO

Could not download old logger – took it with us as we had no batteries to attempt replacement. It later turned out to be nonfunctional, although we kept the cable which works.

Calibration began at 10:25 am. Not recorded by HOBO. Tested HOBO again before leaving and it seemed to be working with test tips.

June 9, 2008

HOBO may or may not have worked during week. No tips recorded. Had trouble again getting it to register tests. Eventually it did successfully record our calibration data.

Old, nonfunctional logger replaced with the one from Gauge 7. New batteries and time changed.

Calibration of 100 mm and 300 mm nozzles successfully completed. We had some thunder and light showers but rigged an umbrella over the tube, and feel that this was successful since there was hardly any wind.