Notes to Francisco (and others) re TB3 Rain Gauges – From Steve Burges

Objective

The goal of the calibration procedure detailed in this protocol is to put a measured quantity of water through each rain gauge so that it can be compared with the quantity that the gauge records.  This is done at more than one (in this case, three) water input rates to estimate the recording bias across a range of rain intensities, and adjustment factors are generated for each input rate.  Without this calibration, the rain gage is likely to have a systematic bias, particularly for high input intensities, due to the imperfect operation of the tipping action and other mechanisms.

Preliminary Steps

Write with a marking pen the gauge number on the underside of each gauge – number 1 to 6.

Record the serial number for each gauge (located on the underneath or inside) and enter gauge number and serial number for each of them on the spread sheets contained within the Excel Workbook entitled:

Mariott Tube Volume Check 7-10-13-2004.xls.

There is a separate spread sheet for each gauge labeled, e.g., TB3 1 Calib, .. TB3 6 Calib.

Mount the two lengths of lumber on the cinder blocks.

Set up the six gauges on the boards - the base of the gauges will be about 40 mm (1.6 in) above the solid floor (possibly on the back or front porch).

The first step of calibration is to run water through each gauge at the maximum rate of approximately 2 m depth of “rain”.

Place a 1.8 Kg Yogurt container under each gauge.

Attach (probably needed) a length of plastic tubing to each of the drain pipes from a TB3 so that the tubing reaches approximately 8 cm below the rim of the container.  This is to ensure catching all the water that is placed through the funnel and reusing it.  (Do not let the tubing hit the bottom, or it may interfere with the flow through the tube).

Place two (2) Mariott calibration tube volumes of water into a large yogurt container (or any suitable container) and pour into the TB3 funnel.  (I think that you will get the full volume in one go.  The TB3 will tip at a constant rate.  The water delivery will be about the equivalent of 500 mm/hr and all of it will e collected into the yogurt collector container below.

Repeat for TB3 number 2 through 6.  This will be a production line job.

As soon as the funnel has emptied, remove the yogurt container – it will have about 2*650 = 1300 ml of water.  Put an empty yogurt container below the TB3 and deliver the water from the ‘full” container to the funnel.  This needs to be done 50 times to generate the equivalent of 2 m of rain.

Keep count of each “iteration” of water fed to the six gauges (there is a form for doing this).

Actual Calibration:

The gauges have been adjusted at the factory to average 0.2 mm/tip,  The adjustment is within a tolerance of +/- 3% at a 100 mm/hr delivery rate.  One gauge could have an actual delivery rate of 0.97*0.2 mm per tip.  Another could have an actual delivery rate of 1.03*0.3 mm/tip and you have no way of knowing this.

Calibrate two gauges at a time.  There are two calibration tubes SJB-1 and SJB–2.  They deliver slightly different volumes that depend on nozzle size (50, 100, and 200 mm/hr nozzles).  The complete calibration details are given on the printed sheets that you use to keep track of how many volumes have been delivered to each rain gauge.

Rain Gauge details:

Load and test the software for downloading the tips from the TB3 gauges.  Set up the data logger (there will be a manual provided) by attaching the terminals to one of the reed switches (there are 2 reed switches on each TB3).  The data logger requires 6 AA batteries.

Launch the logger – add ten or so tips and ensure that you can down load them before starting the long calibration process.

Relaunch the data logger and complete the calibrations (details below).

EXTREMELY IMPORTANT:  	Adjust the time clock for data loggers

The clock on the PC that launches the Hydrologic Services data loggers (and the HOBO) is the clock setting for the data loggers.

Determine if you will use Universal Time (GMT) or local time.

I recommend that you use local time.

This requires setting the PC accurately to local time (PC clocks are terrible and drift about 30 seconds per week.  Timing will be a significant issue for storm details across the rain gauge network.

Calibration Protocol:

At a minimum run five Mariott tube volumes through each TB3 at the 50 mm/hr rate.  Two gauges can be calibrated in 3 hours.  It takes 25 minutes to deliver one tube (about 650 ml) at 50 mm/hr.

N.B.B.  The gauge must be level when calibrating.  Use the plastic wedges that I have provided to level the two gauges that are being calibrated simultaneously.

The 50 mm/hr nozzle (and occasionally the 100 mm/hr nozzle) can be clogged by fine suspended material (or grit).  Use the pins (in the wine corks – This is a good reason to justify pulling a wine cork from time to time.

If time permits, run calibrations at 100 mm/hr and 200 mm/hr.  You can do the 200 mm/h quite fast – about 6 minutes per Mariott tube.

The calibration needs of a TB3 differ from those of a standard tipping bucket gauge.  The TB3 gauge uses a siphon.  8 ml fill the siphon and then are delivered at a constant rate to the tipping bucket mechanism.  If the funnel is full of water the tipping is continuous once it starts.  If the delivery rate is really low, there can be hours between tips.  There are usually small differences in calibrations at the different delivery rates – nowhere nearly as noticeable for the TB3 gauges as for the Texas Industries gauge (see calibration for La Palma gauge)

Gauge Installation – Field sites

The gauge will be supported o a platform that is screwed onto the threaded end of a 2 inch pipe.  Six 3-ft lengths of pipe are available.

Dig a hole 2 feet deep and secure the pipe in it using 0.5 bag of concrete mix, protecting the threaded end of the pipe that attaches to the gauge with either duct tape or a plastic bag taped over it.  Make the concrete “wetter” than usual (a typical ratio of water to concrete is 3/10 or 4/10, but the goal is to create a mixture that is just wet enough to pour and place around the pipe.).  The gauges have to be retrieved in about 2 weeks.  The wet (more water than is good for structural concrete) concrete will have low strength.  This will make it easier to break it free from the pipes when other gauges are removed.  (The pipe and gauges are to be redeployed at another location later this year).

Put the pipe in vertically.  (Use the provided level.  The plate can be screwed in place about 12 hours after the concrete has been first placed.

There are several nuts on each of the studs on the mounting plate that allow leveling the rain gauge.  I have provided wrenches for this purpose.

The rain gauge cowling and funnel is anchored in place by three Allen key machine bolts.  There is a set of Allen keys (wrenches) available for this purpose.

Keep the wrenches and other tools with the gauges when they are being repacked for their next deployment.

Note re Texas Industries Gauge – La Palma.  The rim height is approximately36 inches (about .91 m) above the local ground level.